I’m checking what is actually used for printable Ag/AgCl reference electrodes, especially whether a true one-pot Ag/AgCl ink is practical for inkjet versus making Ag first and chlorinating it afterward. The latter is often cheaper and more robust, so I’m verifying where that tradeoff lands.

The practical answer is looking different from the silver-MOD conductor case: for an Ag/AgCl reference, the cheapest robust route is usually not to jet a chloride-filled conductive ink directly. Most printable references are made by printing silver first, converting only the reference area to Ag/AgCl, then adding a chloride-containing polymer overlayer if you need a more stable solid-state reference.

Thought for 28s

Yes—but for Ag/AgCl reference electrodes, the most practical answer is usually:

Do not try to make a one-bottle inkjettable Ag/AgCl ink first.
Instead, use a 2-step or 3-step process:

- Print silver ink

- Chlorinate the reference area after printing to form Ag/AgCl

- Optionally print or coat a PVB + KCl overlayer if you need a more stable solid-state reference rather than just a pseudo-reference electrode.

That is usually the most reasonable and cost-effective path for the same inkjet/PAD manufacturing environment.

Why this is better than a direct Ag/AgCl ink

A printable Ag/AgCl reference is common in wearable and potentiometric sensors, but the best-performing printed versions are often made by printing Ag and chlorinating it afterward, not by jetting a heavily AgCl-loaded ink. Reviews of printed potentiometric sensors note that the Ag/AgCl layer is commonly followed by a KCl-doped polymer membrane to complete the reference electrode, and fully inkjet-printed solid-state references reported in the literature also used printed Ag first, then chlorination.

That approach is attractive because:

- AgCl is poorly soluble, so a direct “true Ag/AgCl ink” usually becomes a particle-dispersion problem rather than a clean molecular ink problem.

- Particle-based chloride inks are generally harder to filter, less nozzle-friendly, and more variable than printing Ag first and converting only the surface.

- For inkjet, many papers specifically used post-print chlorination: bleach/NaClO, HCl/KCl, or electrochemical chlorination.

So for an inkjet line, Ag first, chlorinate second is usually the practical route.

The most practical manufacturing choices

Option A — cheapest and simplest: pseudo-reference electrode

Use this when you mainly need a local reference for short measurements in a defined electrolyte.

Process

- Print your silver MOD ink or Ag conductive ink.

- Convert the reference zone to Ag/AgCl using:

  - NaClO/bleach, or

  - HCl + KCl, or

  - electrochemical chlorination in KCl.

This is the simplest route and is already demonstrated in inkjet-printed Ag/AgCl references on paper and plastic. One ACS report described a “simple” inkjet-printed Ag/AgCl reference where silver ink was printed and chlorinated with bleach solution. A 2024 inkjet paper-based sensor chlorinated a printed Ag electrode with 0.1 M HCl and 0.1 M KCl.

When to use

- amperometric sensors

- short-run disposable strips

- systems where chloride activity is reasonably controlled

- lowest manufacturing cost

Main weakness

- it behaves more like a pseudo-reference than a classical fully buffered reference, so the potential can drift more if sample chloride and sample composition vary. Reviews note that for potentiometric sensing, a chloride-containing polymer membrane is usually added to improve stability across ionic environments.

Option B — best practical solid-state reference: Ag/AgCl + PVB/KCl overlayer

This is the route I would recommend if you want a real printable reference electrode, not just a pseudo-reference.

Process

- Print silver.

- Chlorinate the reference pad to Ag/AgCl.

- Print or coat a polyvinyl butyral (PVB) + KCl membrane over it.
Some literature variants also include AgNO3 in the membrane, and sometimes a thin Nafion overcoat to reduce salt loss.

Reviews explicitly note that a KCl-doped polymer membrane is commonly printed over the Ag/AgCl layer to complete the reference electrode. The sweat-sensor review also reports that Ag|AgCl|KCl + PVB showed negligible drift over wide pH and chloride ranges and remained stable for about three months.

A representative literature membrane formulation for a flexible solid-state RE used:

- 78.1 mg PVB

- 50 mg KCl

- 50 mg AgNO3

- 1 mL methanol
followed by curing, then 2 µL Nafion as a leakage barrier.

That exact recipe is not automatically the best inkjet formulation, but it strongly supports the underlying design principle: Ag/AgCl base + chloride reservoir in a polymer membrane.

My practical recommendation for your line

For your use case, I would rank the choices like this:

Most cost-effective overall

Print Ag MOD ink → chlorinate locally with NaClO or HCl/KCl

This wins on:

- lowest formulation complexity

- easiest filtration

- least nozzle risk

- no need to keep AgCl particles suspended in the cartridge

- easiest integration into an existing Ag-ink line

Best balance of performance and practicality

Print Ag MOD ink → chlorinate → add a separate PVB/KCl topcoat on the RE area

This is the route to choose for:

- potentiometric sensors

- ion-selective electrodes

- sweat/saliva/environmental sensors where reference stability matters more

Least attractive as a first product

Single-bottle direct Ag/AgCl particle ink for inkjet

I would avoid this first because, for inkjet, it adds:

- dispersion stability problems

- sedimentation risk

- filtration burden

- higher clogging probability

- less process flexibility than post-chlorination

That conclusion is an engineering inference from how the successful inkjet literature is set up: papers repeatedly print silver first and chlorinate afterward, rather than relying on a chloride-particle ink as the main jetting route.

A practical process flow

Here is the version I would actually try in manufacturing:

Route 1: pseudo-reference

Silver ink

- Use the same silver MOD ink platform you are already considering for the conductor.

Print

- Print and dry/cure enough to form a continuous silver film.

Chlorinate

- Mask or selectively wet only the reference area.

- Apply dilute NaClO or HCl/KCl to convert the surface to Ag/AgCl.
Literature examples support both bleach chlorination and HCl/KCl chlorination.

Rinse/dry

- Gentle rinse if needed, then dry.

That is the cheapest route.

Route 2: solid-state reference

Same as above, then add:

Reference membrane overcoat

- PVB dissolved in alcohol, loaded with KCl, optionally a little AgNO3

- Print or microdispense only on the RE area

- Dry/cure gently

- Optional thin Nafion barrier coat to reduce leaching, as used in one flexible saliva-sensor reference design.

Reasonable composition targets

Because the literature gives the architecture more clearly than a universal optimized ink recipe, the practical target should be treated as a formulation window rather than one exact formula.

For the Ag base ink

Use the same kind of particle-free or low-risk Ag ink you would use for your conductive trace. That part should be optimized for jetting, not for chloride chemistry.

For the reference overcoat

A reasonable starting point is:

- PVB as the polymer matrix

- KCl as the chloride reservoir

- optional small AgNO3 addition if you want closer alignment with published solid-state RE formulations

- alcohol solvent system compatible with your substrate and underlying Ag layer

This recommendation is grounded in the literature examples and reviews that explicitly describe KCl-doped polymer membranes, especially PVB + KCl, for stabilizing printed Ag/AgCl references.

Where PAD helps

If your PAD process means plasma-assisted decomposition, then it is very useful for the silver printing step, but it is not the main reason to use a direct Ag/AgCl ink. The reference function still mainly depends on:

- forming an Ag/AgCl interface

- maintaining chloride availability near that interface

So even with PAD, I would still prefer:
Ag MOD printing + PAD conversion + local chlorination + optional PVB/KCl overcoat
over trying to invent a direct jettable Ag/AgCl dispersion ink.

Bottom line

For the same manufacturing platform, the most practical Ag/AgCl strategy is:

Best low-cost route:
print Ag → chlorinate the reference area

Best practical solid-state route:
print Ag → chlorinate → add PVB/KCl overcoat

I would not start with a direct one-bottle Ag/AgCl inkjet ink unless you have a strong reason to accept the added dispersion and nozzle-control burden. The published inkjet RE work and recent reviews strongly favor the post-chlorination architecture instead.